**Placement of Bridge on DRY Pediatric Circuit**

Equipment: ¼” bridge tubing with 1 stopcock; non-venting cap

Steps to insert bridge (with sterile hat, mask, and gloves):

1. Open ¼” bridge tubing (should be located on shelf under Cardio Help).
2. Bridge comes with 2 stopcocks, remove one and leave one in place.
3. Place non-venting cap on side port of stopcock.
4. On circuit venous luer, move pigtail with syringe from end port to side port of stopcock.
5. Attach non-stopcock end of bridge to end port of stopcock on venous luer.
6. Remove non-venting cap from arterial luer.
7. Attach stopcock end of bridge to arterial luer.
8. Turn stopcocks at both ends of bridge off to circuit.